## Slide 1: Progress on Thermal Studies of Mockup Modules
Arthur Campello, Neil Minet, Jose Monroy, Yadira Bordlemay Padilla (Cornell University: Testing)
Matthew Jones, Andy Jung, Souvik Das (Purdue University: Mockup production)
March 14, 2019

## Slide 2
March 14, 2019 A. Campello 1
Aim and Scope of Studies
Build and validate experimental procedure for thermal testing of mockup modules (with and without CO2 cooling)
Determining the relationship between supplied power, minimum, maximum and average temperatures
Comparing the effects of cooling materials and structural application to regulate heat on a module

## Slide 3
Materials and Methods
Mockup-modules were made by Purdue (Mathew Jones), wire bonds sealed at Cornell
Thermal measurements performed using FLIR E50 infrared camera (0.1C precision and +/- 0.13C measured accuracy, last calibrate 2014)
Thermal analysis performed using FLIR Tools® software
Power supplied to modules using KEYSIGHT U8001 single output DC power supply
March 14, 2019 A. Campello 2

## Slide 4
Verification of Camera Accuracy
Camera readings verified against thermocouple reading of water in thermal equilibrium
Performed for eight measurements between 0ºC and 25ºC
Difference between TC and camera:
  Mean: -0.04ºC, STD: 0.08ºC
Camera was calibrated in 2012 and 2014, FLIR claims +/-2ºC accuracy after five years
16.4
Fan
TC
Stir
rod
Cup
March 14, 2019 A. Campello 3

## Slide 5
Thermal Analysis Software
Thermal analysis software reports
Maximum, minimum, and average Temp and temps. At any point of the 240x180 size infrared image
March 14, 2019 A. Campello 4

## Slide 6
IR Maps of Front and Back of Module Suspended in Air
Both images taken at 2.5W for 40 minutes from 4 inches using IR camera
Modules positioned vertically
Temperature comparisons:
 | Front | Back
Minimum | 26.1ºC | 46.1ºC
Maximum | 64.0ºC | 67.3ºC
Average | 48.5ºC | 53.7ºC
Front
Back
March 14, 2019 A. Campello 5

## Slide 7
Temperature over Power of Module Front Suspended in Air
Min, max, and avg T of the front of one module, recorded three times for 0.5, 1, 1.5, 2, and 2.5 watts
Error bars show 95% confidence intervals
ΔT/ΔW Slopes (Y=mx+b) allow for temperature extrapolation
  Min: 0.96
  Max: 15.92
  Avg: 9.94
March 14, 2019 A. Campello 6

## Slide 8
Heat Dissipation Study Approaches
Developing a mechanism to secure a module (and a backing) to a heat sink (aluminum block)
IR analysis can measure steady-state T in different heat dissipation configs
Temperature to be measured with RTD array to provide quantifiable results
Will compare with simulations
March 14, 2019 A. Campello 7

## Slide 9
Future Steps
Further thermal measurements of module in air to yield a complete thermal profile
Development of a thermal simulation for the mockup modules
Development of a heat dissipation test-bed to cool module using CO2
Measurements of heat dissipation through various backing materials
March 14, 2019 A. Campello 8